## Sheet1
Text
Based on the detailed Software Requirements Specification provided, your interpretation that the GUI may be in the state "Start" or "SelfTest" or even in the state "StartUp" when the controller is in the state "SelfTest" is not directly supported by the specifications given. The specifications outline specific transitions between states for both the GUI and the controller, but they do not explicitly state that the GUI can be in the "StartUp" state while the controller is in the "SelfTest" state.
Here's a breakdown based on the provided specifications:
1. **Start-up State (GUI and Controller)**: Both the GUI and the controller have a "Start-up" state. The GUI transitions from "Start-up" to either "Start" or "Ventilation" based on whether ventilation is on or off. The controller transitions from "Start-up" to "SelfTest" after completing its start-up checks.
2. **SelfTest State (GUI and Controller)**:
- For the GUI, the "SelfTest" state allows the user to perform a sequence of tests, and upon passing these tests, it transitions to the "Menu" state.
- For the controller, after completing the self-test, it transitions to the "VentilationOff" state if the self-test is passed or if the GUI asks for resuming ventilation.
3. **GUI in Start or SelfTest State**:
- The GUI can indeed be in the "SelfTest" state as per GUI.1.4 and GUI.9, which describe the transition from "SelfTest" to "Menu" upon passing the self-test.
- The GUI can be in the "Start" state as per GUI.1.2, which allows the user to resume ventilation or start ventilation for a new patient. This is a logical state for the GUI to be in before or after the "SelfTest" state depending on user actions.
4. **GUI in StartUp State**:
- The specifications do not explicitly support the GUI being in the "StartUp" state while the controller is in the "SelfTest" state. The "StartUp" state for the GUI is primarily for initialization with default configuration parameters and determining whether to transition to "Start" or "Ventilation" based on the ventilation status.
Therefore, while the GUI can logically be in the "Start" or "SelfTest" states when the controller is in the "SelfTest" state, the specifications do not support the GUI being in the "StartUp" state at this point. The transition paths outlined in the specifications suggest a sequential flow where the GUI would move out of "StartUp" before the controller enters "SelfTest".